## Terminal Blocks
PFS Terminal Block wiring identification | [7 empty] | 2018-04-01
TB1 - Inside the Hoffman box to the left of the EMCO power supply
Device | Terminal | Wire | TB1 | TB1 | Wire | Terminal | Device | Notes
AC plug | Neutral | Blu | 1 | 2 | Blu | 5 | TB3
Temperature controller | Neutral | Wht | 1 | 2 | Blu |  | Flowmeter totalizer
AC plug | Line | Brn | 3 | 4 | Brn | 7 | TB3
Temperature controller | Line | Blk | 3 | 4 | Brn |  | Flowmeter totalizer
TB2 - Inside Hoffman box to the right of the EMCO power supply
Device |  | Wire | TB2 | TB2 | Wire |  | Device | Notes
|  |  | 1 | 2 | Blk 18 | 27 | TB5 | Return
EMCO HC2024R |  | Blk & Blk | 1 | 2 | Blu | 6 | TB19 | Return line for ThAr lamp
EMCO HC2024R |  | Green | 3 | 4 | Gry/Blk | 45 | TB5 | To pot and F2-08AD-2 ana 1,5
EMCO HC2024R |  | Red thin | 5 | 6 | Yel 18 | 8 | TB4 | 24V input
EMCO HC2024R |  | Yel | 7 | 8 | Yel/Gry | 37 | TB5 | To pull-down and F2-08TRS relay 0,6
EMCO HC2024R |  | Red thick | 9 | 10 | Vio | 8 | TB19 | Hot line for ThAr lamp
TB3 - Inside Hoffman box beneath power supplies
Device | Terminal | Wire | TB3 | TB3 | Wire | Terminal | Device | Notes
SMC100 power supply |  | Wht | 1 | 2 | Blu | AC Input | PLC Base 1
TB3 | 3 | Jump | 1 | 2 | Blu | 4 | TB11
TB3 | 1 | Jump | 3 | 4
TB3 | 5 | Jump | 3 | 4 | Blu | 4 | HBB512-A 5/12V supply
TB3 | 3 | Jump | 5 | 6 | Blu | 4 | SLS-05-060-1T 5V supply
TB1 | 2 | Blu | 5 | 6 | Blu | 4 | HC24-2.4-A 24V supply
TB1 | 4 | Brn | 7 | 8 | Brn | AC Input | PLC Base 1
TB3 | 9 | Jump | 7 | 8 | Brn | 1 | TB11
TB3 | 7 | Jump | 9 | 10
TB3 | 11 | Jump | 9 | 10 | Brn | 1 | HBB512-A 5/12V supply
TB3 | 9 | Jump | 11 | 12 | Brn | 1 | SLS-05-060-1T 5V supply
SMC100 power supply |  | Blk | 11 | 12 | Brn | 1 | HC24-2.4-A 24V supply
TB4 - Inside Hoffman box beneath power supplies
Device | Terminal | Wire | TB4 | TB4 | Wire | Terminal | Device | Notes
F2-08TRS Relay - base 0, slot 6 | NO1 | Red | 1 | 2 | Org/Red | 42 | TB6 | +12V
H2-CTRIO High speed Counter | 1M | Org | 1 | 2 | Org 18 | +OUT | HBB512-A 5/12V supply | +12V
TB7 | 1 | Red 18 | 3 | 4 | Red 18 | 1 | Board connector | +5V from Power One supply
F2-08TRS relay - base0, slot6 | C4 | Red 18 & Red | 3 | 4 | Red 18 | +5VDC | HBB512-A 5/12V supply | +5V from Power One supply
|  |  | 5 | 6
|  |  | 5 | 6
Board 2 | Y | Yel | 7 | 8 | Yel 18 | 6 | TB2 | +24V
TB8 | 1 | Yel | 7 | 8 | Jump | 10 | TB4 | +24V
TB5 | 23 | Yel 18 | 9 | 10 | Jump | 8 | TB4 | +24V
Board connector | 54 | Yel/Grn | 9 | 10 | Jump | 12 | TB4 | +24V
Dewar vacuum controller |  | Gry | 11 | 12 | Jump | 10 | TB4 | +24V
Grating vacuum controller |  | Gry | 11 | 12 | Yel 18 | +OUT | HC24-2.4-A 24V supply | +24V
Dewar vacuum controller |  | Gry | 13 | 14 | Black | 16 | TB6 | Common
Grating vacuum controller |  | Gry | 13 | 14 | Jump | 16 | TB4 | Common
STP-DRV-4035 | VDC- | Bk 18 | 15 | 16 | Jump | 14 | TB4 | Common
TB5 | 27 | Blk 18 | 15 | 16 | Jump | 18 | TB4 | Common
TB7 | 3 | Blk | 17 | 18 | Jump | 16 | TB4 | Common
TB4 | 19 | Jump | 17 | 18 | Blk 18 | 12 | TB8 | Common
TB4 | 17 | Jump | 19 | 20 | Blk 18 | 5V RTN | HBB512-A 5/12V supply | Common
TB4 | 21 | Jump | 19 | 20 | Blk 18 | -OUT | HBB512-A 5/12V supply | Common
TB4 | 19 | Jump | 21 | 22 | Blk 18 | -OUT | SLS-05-060-1T 5V supply | Common
Connector J5 | P & E | Blk & Blu | 21 | 22 | Blk 18 | -OUT | HC24-2.4-A 24V supply | Bar. Press. Sensor Grounds
TB6 | 33 | Wht 18 | 23 | 24 | Red 18 | +OUT | SLS-05-060-1T 5V supply | +5V from Sola supply (for lamp)
|  |  | 23 | 24 |  |  |  | +5V from Sola supply (for lamp)
Board connector | 53 | Grn/Red | 25 | 26 | Grn/Red | 10 | TB6 | +9V
|  |  | 25 | 26 |  |  |  | +9V
|  |  | 27 | 28
|  |  | 27 | 28
Grating vacuum controller |  | Brn | 29 | 30 |  |  |  | Ground
Dewar vacuum controller |  | Brn | 29 | 30 | Grn 18 |  | Electronics mounting plate | Ground
TB5 - Inside Hoffman box near right PLC base
Device | Terminal | Wire | TB5 | TB5 | Wire | Terminal | Device | Notes
ZipLink ZL-CM40 terminal block | CIII low | Red | 1 | 2 | Red 18 | 24 | TB6 | +5V
TB4 | 3 | Red 18 | 1 | 2 | Jump | 4 | TB5 | +5V
ThAr brightness pot |  | Red/Org | 3 | 4 | Jump | 2 | TB5 | +5V
Slit Linear Transducer |  | Red | 3 | 4 | Jump | 6 | TB5 | +5V
Hartmann Rotary Transducer |  | Yel | 5 | 6 | Jump | 4 | TB5 | +5V
Focus Rotary Transducer |  | Blu | 5 | 6 | Jump | 8 | TB5 | +5V
Iodine linear transducer |  | Red | 7 | 8 | Jump | 6 | TB5 | +5V
|  |  | 7 | 8 | Red | C | Connector J5 | Bar. Pressure. Sensor Trigger +5V
To board, 1K Ohmite L12J1K0E | 51 | Red | 9 | 10 | Red | 0 | AC output module F2-08TA
To focus motor black |  | Red | 9 | 10
To board, 2uF Illinois PHC400VW2 | 52 | Blk | 11 | 12 | Blk | 1 | AC output module F2-08TA
To focus motor red |  | Blk | 11 | 12
To focus motor white |  | Wht | 13 | 14 | Blue 18 | AC Input | PLC Base 2
Ion pump interlock pin 7 |  | Wht | 15 | 16 | Wht | C3 | F2-08TRS relay - base 0, slot 6
Ion pump interlock pin 3 |  | Grn | 17 | 18 | Grn | NO3 | F2-08TRS relay - base 0, slot 6
ThAr brightness pot |  | Blk/Gry | 19 | 20 | Resistor | 26 | TB5 | 4.22K
F2-08AD-2 ana in - base 1, slot 7 | +24V | Yel/Org | 21 | 22 | Yel 18 | C5 | F2-08TRS relay - base 0, slot 4 | +24V
TB5 | 25 | Jump | 21 | 22 | Yel/Org | +24V | F2-08DA2 ana out - base 1,slot 0
TB5 | 23 | Jump | 23 | 24 | Yel/Brn | +24V | F2-08AD-2 ana in - base 1, slot 3
TB4 | 9 | Yel 18 | 23 | 24 | Yel/Brn | +24V | F2-08AD-2 ana in - base 1, slot 4
Iodine linear transducer |  | Blk | 25 | 26 | Resist, Grn | 20,54 | TB5 | Common
TB5 | 50 | Blk | 25 | 26 | Jump | 28 | TB5 | Common
TB4 | 15 | Blk 18 | 27 | 28 | Jump | 26 | TB5 | Common
TB2 | 2 | Blk 18 | 27 | 28 | Jump | 30 | TB5 | Common
Slit linear transducer |  | Blk | 29 | 30 | Jump | 28 | TB5 | Common
TB5 | 31 | Jump | 29 | 30 | Blk | 0V | F2-08DA2 ana out - base 1,slot 0 | Common
TB5 | 29 | Jump | 31 | 32 | Blk | 0V | F2-08AD-2 ana in - base 1, slot 5 | Common
TB5 | 33 | Jump | 31 | 32 | Blk | 0V | F2-08AD-2 ana in - base 1, slot 3 | Common
TB5 | 31 | Jump | 33 | 34 | Blk | 42 | TB5 | Common
Hartmann Rotary Transducer |  | Grn | 33 | 34 | Jump | 34 | TB5 | Common
Focus Rotary Transducer |  | Blk | 35 | 36 | Jump | 32 | TB5 | Common
TB5 | 37 | Resistor | 35 | 36 | Blk | 0V | F2-08AD-2 ana in - base 1, slot 4 | Common
TB5 | 35 | Resistor | 37 | 38 | Yel/Brn | NO2 | F2-08TRS relay - base 0, slot 4
TB2 | 8 | Yel/Gry | 37 | 38
Slit Linear Transducer |  | Org | 39 | 40 | Brn/Yel | CH1+ | Analog 8CH - base 1, slot 3 | Transducer Orange
1uF capacitor | 41 | Capacitor | 39 | 40
1uF capacitor (2) | 39/43 | Capacitor | 41 | 42 | Blk | 34 | TB5 | to common
1uF capacitor | 41 | Capacitor | 43 | 44
Focus Rotary Transducer |  | Wht | 43 | 44 | Yel/Brn | CH2+ | Analog 8CH - base 1, slot 3 | Transducer Orange
TB2 | 4 | Gry/Blk | 45 | 46 | Gry/Blk | CH3+ | F2-08AD-2 ana in - base 1, slot 3
ThAr brightness pot | middle | Gry/Blk | 45 | 46 |  |  |  | *need to check pot value
Iodine linear transducer |  | Org | 47 | 48 | Yel/Brn | CH7+ | Analog 8CH - base 1, slot 3 | Transducer orange
1uF capacitor | 49 | Capacitor | 47 | 48
1uF capacitor (2) | 47/51 | Capacitor | 49 | 50 | Blk | 25 | TB5 | to common
1uF capacitor | 49 | Capacitor | 51 | 52
Hartmann Rotary Transducer |  | Org | 51 | 52 | Vio/Gry | CH8+ | Analog 8CH - base 1, slot 3 | Transducer Orange
CCD STAT- (charge shuffle) |  | Grn | 53 | 54 | Grn | 26 | TB5 | Common
CCD STAT+ (charge shuffle) |  | Wht | 55 | 56 | Wht | +V1 | F2-08DA2 ana out - base 1,slot 0 | +5V trigger for CCD charge shuffle
TB6 - Inside Hoffman box near left PLC base
Device | Terminal | Wire | TB6 | TB6 | Wire | Terminal | Device | Notes
Flipper control pad | TO | Wht | 1 | 2 | Wht |  | Flipper | Thru connector, TB
Flipper control pad | TO | Red | 3 | 4 | Yel 18 |  | Flipper | Thru connector, TB
Flipper control pad | TO | Shield | 5 | 6 | Grn |  | Flipper | Thru connector, TB
Flipper control pad | TTL | Grn/Yel | 7 | 8 | Grn/Yel | NO1 | F2-08TRS relay - base 0, slot 4
Flipper control pad | DC | Grn/Red | 9 | 10 | Grn/Red | 26 | TB4 | 9V
STP-DRV-4035 | STEP- | Wht/Org | 11 | 12 | Wht/Org | 3 | DT-CTRINT
STP-DRV-4035 | DIR- | Org/Wht | 13 | 14 | Org/Wht | 4 | DT-CTRINT
Flipper control pad | TTL | Yel/Grn | 15 | 16 | BLk | C4 | DT-CTRINT | Ground
TB6 | 21 | Jump | 15 | 16 | Blk | 14 | TB4 | Ground
TB6 | 19 | Jump | 17 | 18 | Blk |  | TB25 | Ground for air heater fan
TB6 | 23 | Jump | 17 | 18 | Blk/Org | NC6,NC6 | F2-08TRS relays - base 0, slot 3,4
TB6 | 21 | Jump | 19 | 20 | Blk/Gry | NC6 | F2-08TRS relay - base 0, slot 5 | Common
|  |  | 19 | 20 | Jump | 26 | TB6 | Common
Flipper control pad | DC | Red/Grn | 21 | 22 | Jump | 24 | TB6 | Common
SMC100 | GPIO 15 | Blk | 21 | 22 | Blk | 28 | TB6 | Common
STP-DRV-4035 | STEP+/DIR+ | Red | 23 | 24 | Red | C2 | DT-CTRINT | +5V
TB6 | 29 | Jump | 23 | 24 | Red 18 | 2 | TB5 | +5V
TB6 | 27 | Jump | 25 | 26 | Red | C2 | F2-08TRS relay - base 0, slot 4 | +5V
|  |  | 25 | 26 | Org/Blk | C1 | F2-08TRS relay - base 0, slot 4 | +5V
Common for Iodine limit switches |  | Wht & Blk | 27 | 28 | Blk | 22 | TB6 | Common
QTH ground |  | Blk | 27 | 28 | Blk/Gry | NC6 | F2-08TRS relay - base 0, slot 6 | Common
STP-DRV-4035 | VDC+ | Yel 18 | 29 | 30 | Yel 18 | NO5 | F2-08TRS relay - base 0, slot 4 | +24V
Slit speed pot |  | Vio/Org | 31 | 32 | Org/Blk | C0 | F2-08TRS relay - base 0, slot 3 | * need to check pot value
TB4 | 23 | Wht 18 | 33 | 34 | Wht 18 | C3 | F2-08TRS relay - base 0, slot 3 | 5V from Sola power supply
|  |  | 33 | 34
Iodine speed pot |  | Gry/Red | 35 | 36 | Red | NO3 | F2-08TRS relay - base 0, slot 4 | * need to check pot value
|  |  | 35 | 36 | Org/Red | C0 | F2-08TRS relay - base 0, slot 4
22 Ohm |  | Resistor | 37 | 38 | Red | C3 | F2-08TRS relay - base 0, slot 4
Quartz lamp |  | Red | 39 | 40 | Red 18 | NO3 | F2-08TRS relay - base 0, slot 3
Iodine speed pot |  | Red/Gry | 41 | 42 | Org/Red | 2 | TB4 | to 12V supply
22 Ohm |  | Resistor | 41 | 42 | Jump | 44 | TB6
Slit speed pot |  | Org/Vio | 43 | 44 | Jump | 42 | TB6
Connector J5 | R | Grn | 43 | 44 | Org/Red | 64 | TB6 | Bar. Press. Supply +12VDC
Slit Motor Brown |  | Brn | 45 | 46 | Brn/Vio | C6 | F2-08TRS relay - base 0, slot 3
Slit Motor Gray |  | Wht | 47 | 48 | Vio/Brn | C7 | F2-08TRS relay - base 0, slot 3
Iodine motor brown |  | Brn | 49 | 50 | Blu/Vio | C6 | F2-08TRS relay - base 0, slot 4
Iodine motor gray |  | Blu | 51 | 52 | Vio/Blk | C7 | F2-08TRS relay - base 0, slot 4
Shutter gray |  | Blk | 53 | 54 | Brn/Vio | C6 | F2-08TRS relay - base 0, slot 5
Shutter brown |  | Red | 55 | 56 | Vio/Brn | C7 | F2-08TRS relay - base 0, slot 5
Diffuser shutter gray |  | Yel | 57 | 58 | Blu/Vio | C6 | F2-08TRS relay - base 0, slot 6
Diffuser shutter brown |  | Org | 59 | 60 | Vio/Blk | C7 | F2-08TRS relay - base 0, slot 6
101 ohm | 63 | Resistor | 61 | 62 | Org/Vio | C0 | F2-08TRS relay - base 0, slot 5 | for shutter
101 ohm and 751 ohm | 61/63 | Resistor (2) | 63 | 64 | Org/Red | 44 | TB6 | 12V
751 ohm | 63 | Resistor | 65 | 66 | Org/Vio | C0 | F2-08TRS relay - base 0, slot 6 | for diffuser shutter
TB7 - Inside Hoffman box, right of board
Device | Terminal | Wire | TB7 | TB7 | Wire | Terminal | Device | Notes
F2-08TRS - base 0, slot 6 | NO4 | Red/Wht | 1 | 2 | Red | A | Board 2 | +5V
|  |  | 1 | 2 | Grn |  | PMT red via cable 11 | +5V
TB4 | 17 | Blk | 3 | 4 | Blk | Z | Board 2 | Common
|  |  | 3 | 4 | Yel |  | PMT black | Common
F2-CP128 Port 3 |  | Grn | 5 | 6 | Org |  | PMT serial (DB-9 pin 3)
F2-CP128 Port 3 |  | Red | 7 | 8 | Red |  | PMT serial (DB-9 pin 2)
F2-CP128 Port 3 |  | Wht & Blu | 9 | 10 | Brn |  | PMT serial (DB-9 pin 5)
0-250 Ohm pot for PMT heater |  | Brn | 11 | 12 | Brn |  | 250 Ohm resistor by PMT (heat) | 20VAC neutral: 0.49W desired = 0.044A
F2-08TRS Relay - base 0, slot 5 | NO3 | Blu | 13 | 14 | Red |  | 250 Ohm resistor by PMT (heat) | 20VAC hot: 0.49W desired = 0.044A
SMC100 motor cable |  | Wht 30 | 15 | 16 | Wht |  | Newport LTA-HL actuator (CCD) | SMC100 motor cable is actually 2
Ziplink ZL-CM40 | CII lo | Red | 17 | 18 | Brn |  | Newport LTA-HL actuator (CCD) | cables inside the black insulation.
SMC100 motor cable |  | Brn 30 | 17 | 18 |  |  |  | The DB25 connector has an electronic
SMC100 motor cable |  | Blk 30 | 19 | 20 | Blk |  | Newport LTA-HL actuator (CCD) | board attached (soldered), so only 12
SMC100 motor cable |  | Org 30 | 21 | 22 | Org |  | Newport LTA-HL actuator (CCD) | wires come out through the cable even
SMC100 motor cable |  | Blu 30 | 23 | 24 | Blu |  | Newport LTA-HL actuator (CCD) | though 17 of the DB25 pins are
SMC100 motor cable |  | Grn 28 | 25 | 26 | Grn |  | Newport LTA-HL actuator (CCD) | connected. Cable groups:
SMC100 motor cable |  | Gry 30 | 27 | 28 | Yel |  | Newport LTA-HL actuator (CCD) | 1. wht, blk, org, grn
SMC100 motor cable |  | Red 30 | 29 | 30 | Red |  | Newport LTA-HL actuator (CCD) | 2. wht, blk, blu, brn, rd, org, gry
SMC100 motor cable |  | Org 28 | 31 | 32 | Org |  | Newport LTA-HL actuator (CCD) | 3. shield
SMC100 motor cable |  | Brn 28 | 33 | 34 | Brn |  | Newport LTA-HL actuator (CCD)
SMC100 motor cable |  | Wht 28 | 35 | 36 | Wht |  | Newport LTA-HL actuator (CCD)
SMC100 motor cable |  | Shield | 37 | 38 | Shield |  | Newport LTA-HL actuator (CCD)
TB8 - Inside Hoffman box, below board
Device | Terminal | Wire | TB8 | TB8 | Wire | Terminal | Device | Notes
F2-08AD-2 ana in, I2 cell T sensor | CH1+ | Org & Org | 1 | 2 | Resistor | 4 | TB8
TB4 |  | Yel | 3 | 4 | Resistor | 2 | TB8 | +24V
Flowmeter totalizer |  | Vio/Grn | 5 | 6 | Vio/Grn | 45 | Board 1
Grating vacuum gauge controller |  | Yel/Grn | 7 | 8 | Ylw/Grn | 47 | Board 1
Dewar vacuum gauge controller |  | Blk/Grn | 9 | 10 | Blk/Grn | 49 | Board 1
Dewar vacuum gauge controller |  | Grn/Blk | 11 | 12 | Blk 18 | 18 | TB4 | Common
Grating vacuum gauge controller |  | Grn/Yel | 11 | 12 | Jump | 16 | TB8 | Common
Flowmeter totalizer |  | Grn/Pur | 13 | 14 | Jump | 14 | TB8 | Common
|  |  | 13 | 14 | Blk 18 | 55 | Board 1 | Common
Thermistor T1 (table-grating) |  | Blk | 15 | 16 | Blk/Brn | 31 | Board 1 | Common
Thermistor T2 (table-middle) |  | Grn | 17 | 18 | Brn/Red | 32 | Board 1 | Common
Thermistor T3 (table-CCD) |  | Blu | 19 | 20 | Gry/Yel | 33 | Board 1 | Common
Thermistor T4 (air-inside) |  | Blu | 21 | 22 | Blk/Grn | 34 | Board 1 | Common
Thermistor T5 (air-outside) |  | Red | 23 | 24 | Brn/Blk | 35 | Board 1 | Common
Thermistor T6 (dewar shell) |  | Brn | 25 | 26 | Brn/Vio | 36 | Board 1 | Common
Thermistor T7 (glycol supply) |  | Org | 27 | 28 | Brn/Yel | 37 | Board 1 | Common
Thermistor T8 (glycol return) |  | Blk | 29 | 30 | Blk/Org | 38 | Board 1 | Common
Thermistor T1 (table-grating) |  | Wht | 31 | 32 | Wht | 27 | Board 1 | Rset
Thermistor T1 (table-grating) |  | Blu | 33 | 34 | Blu | 26 | Board 1 | Thermistor
Thermistor T2 (table-middle) |  | Brn | 35 | 36 | Brn | 25 | Board 1 | Rset
Thermistor T2 (table-middle) |  | Grn | 37 | 38 | Grn | 24 | Board 1 | Thermistor
Thermistor T3 (table-CCD) |  | Yel | 39 | 40 | Yel | 21 | Board 1 | Rset
Thermistor T3 (table-CCD) |  | Grn | 41 | 42 | Grn | 20 | Board 1 | Thermistor
Thermistor T4 (air-inside) |  | Yel | 43 | 44 | Yel | 18 | Board 1 | Rset
Thermistor T4 (air-inside) |  | Wht | 45 | 46 | Wht | 17 | Board 1 | Thermistor
Thermistor T5 (air-outside) |  | Wht | 47 | 48 | Wht | 15 | Board 1 | Rset
Thermistor T5 (air-outside) |  | Blk | 49 | 50 | Blk | 14 | Board 1 | Thermistor
Thermistor T6 (dewar shell) |  | Yel | 51 | 52 | Yel | 12 | Board 1 | Rset
Thermistor T6 (dewar shell) |  | Wht | 53 | 54 | Wht | 11 | Board 1 | Thermistor
Thermistor T7 (glycol supply) |  | Yel | 55 | 56 | Yel | 9 | Board 1
Thermistor T7 (glycol supply) |  | Blu | 57 | 58 | Blu | 8 | Board 1
Thermistor T8 (glycol return) |  | Grn | 59 | 60 | Grn | 6 | Board 1
Thermistor T8 (glycol return) |  | Wht | 61 | 62 | Wht | 5 | Board 1
TB9 - Inside Hoffman box, below and to right of board
Device | Terminal | Wire | TB9 | TB9 | Wire | Terminal | Device | Notes
|  |  | 1 | 2 | Wht/Grn |  | Board
Ion pump controller analog out | DB9 pin2 | Red | 3 | 4 | Resistor | 6 | TB9 | 10K resistor for V divider
PLC Analog input, Base 1, Slot 3 | Ch4+ | Wht/Brn | 5 | 6 | Resistor | 4,8 | TB9 | 10K resistor for V divider
Ion pump controller common | DB9 pin 6 | Blk | 7 | 8 | Resistor | 6 | TB9 | 10K resistor for V divider
Cable 2 |  | Brn | 9 | 10 | Org/Ylw | W | Board 2 | Guider shutter
F2-08TRS Relay - base 0, slot 5 | NO4 | Blu | 11 | 12 | Grn/Yel | X | Board 2
Cable 2 |  | Blu | 11 | 12 |  |  |  | Guider shutter
RG174U coaxial cable to PMT |  | Blk | 13 | 14 | 16 | TB9 | 56 ohm 2W resistor
|  |  | 13 | 14 | Blu/Yel | B | Board 2 | TTL output from PMT
RG174U coaxial cable to PMT |  | Shield | 15 | 16 | Blk | 3 | TB7 | Ground
|  |  | 15 | 16 | 14 | TB9 | 56 ohm 2W resistor
|  |  | 17 | 18
|  |  | 19 | 20
|  |  | 21 | 22
|  |  | 23 | 24
Cable to dewar mount: rear heater |  | Red | 25 | 26 | Red & Org | 19 | TB11 | 0-24VDC to TB23-5 for rear dewar heat
Cable to dewar mount: rear heater |  | Blk | 27 | 28 | Blk & Grn | 17 | TB11 | Ground to TB23-7 for rear dewar heat
Cable to dewar mount: lN2 fill heater |  | Black | 29 | 30 |  |  | 0-25 Ohm pot for lN2 fill heater | Neutral, 16VAC nominal, adjusted
Cable to dewar mount: side heater |  | Black | 31 | 32 | Brn |  | 0-25 Ohm pot for dewar side heater | Neutral, 16VAC nominal, adjusted
0-250 Ohm pot for PMT heater |  | Brn | 33 | 34 | Brn 18 |  | FD5-16 Transformer | Neutral, 16VAC nominal
0-25 Ohm pot for lN2 fill heater |  |  | 33 | 34 | Brn |  | 0-25 Ohm pot for dewar side heater | Neutral, 16VAC nominal
F2-08TRS Relay - base 0, slot 5 | C3 | Blu | 35 | 36 | Blu 18 |  | FD5-16 Transformer | Hot, 16VAC nominal
|  |  | 35 | 36
Cable to dewar mount: side heater |  | Red | 37 | 38 | Blu | NO5 | F2-08TRS Relay - base 0, slot 5 | Hot, 16VAC nominal
Cable to dewar mount: lN2 fill heater |  | White | 37 | 38
|  |  | 39 | 40
|  |  | 39 | 40
TB10 - Inside Hoffman box, lower right next to power supplies
Device | Terminal | Wire | TB10 | TB10 | Wire | Terminal | Device | Notes
|  |  | 1 | 2 |  |  |  | Ground
Shield for shielded cable |  | Blk | 1 | 2 | Jump | 4 | TB10 | Ground for pre-slit assembly
Ground for thermal fluid pipes |  | Grn | 3 | 4 | Jump | 2 | TB10 | Ground
Shield for shielded cable |  | Shield | 3 | 4 | Jump | 6 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 5 | 6 | Jump | 4 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 5 | 6 | Jump | 8 | TB10 | Ground
Shield for shielded cable |  | Shield | 7 | 8 | Jump | 6 | TB10 | Ground
Shield for shielded cable |  | Shield | 7 | 8 | Jump | 10 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 9 | 10 | Jump | 8 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 9 | 10 | Jump | 12 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 11 | 12 | Jump | 10 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 11 | 12 | Jump | 14 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 13 | 14 | Jump | 12 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 13 | 14 | Jump | 16 | TB10 | Ground
Flowmeter totalizer |  | Grn 18 | 15 | 16 | Jump | 14 | TB10 | Ground for flowmeter totalizer
Shield for shielded cable |  | Shield | 15 | 16 | Jump | 18 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 17 | 18 | Jump | 16 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 17 | 18 | Jump | 20 | TB10 | Ground for shielded cable
|  |  | 19 | 20 | Jump | 18 | TB10 | Ground
|  |  | 19 | 20 | Jump | 22 | TB10 | Ground
|  |  | 21 | 22 | Jump | 20 | TB10 | Ground
Shield for shielded cable |  | Shield | 21 | 22 | Jump | 24 | TB10 | Ground for shielded cable
Shield for shielded cable |  | Shield | 23 | 24 | Jump | 22 | TB10 | Ground for shielded cable
Flowmeter totalizer |  | Grn 18 | 23 | 24 | Grn 18 |  | Electronics mounting plate | Ground
TB11 - Inside Hoffman box, middle right next to power supplies
Device | Terminal | Wire | TB11 | TB11 | Wire | Terminal | Device | Notes
TB3 | 8 | Brn 18 | 1 | 2 | Brn 18 | AC Input | Modified HD12-6.8-AG
AC output module F2-08TA, & Transformer | C4-7 | Brn 18 | 1 | 2 | Brn 18 | AC Input | Modified HD24-4.8-AG
AC Heat tape, air heat resistors 1 & 2 |  | Blu18,Blu,Wht | 3 | 4 | Blu 18 |  | Transformer | Blu18-heat tape, Blu+Wht air heat
TB11 | 5 | Jump | 3 | 4 | Blu 18 | 2 | TB3
TB11 | 3 | Jump | 5 | 6 | Blu 18 | AC Input | Modified HD12-6.8-AG
AC Heat tape |  | Wht | 5 | 6 | Blu 18 | AC Input | Modified HD24-4.8-AG
15W AC Heat tape |  | Brn | 7 | 8 | Brn | 4 | AC output module F2-08TA
30W AC Heat tape |  | Red | 9 | 10 | Red | 5 | AC output module F2-08TA
45W AC Heat tape |  | Grn | 11 | 12 | Grn | 6 | AC output module F2-08TA
60W AC Heat tape |  | Blk | 13 | 14 | Blk | 7 | AC output module F2-08TA
|  |  | 15 | 16 |  |  |  | Used as buffer between AC and DC
TB9 | 38 | Blu & Grn | 17 | 18 | Blk 18 |  | Modified HD24-4.8-AG | Ground - for side dewar heat tape
TB9 | 40 | Red & Org | 19 | 20 | Yel 18 |  | Modified HD24-4.8-AG | +24VDC - for side dewar heat tape
Water pump |  | Blk 14 | 21 | 22 | Blk 18 |  | Modified HD12-6.8-AG | Ground
Water pump |  | Red 14 | 23 | 24 | Org 18 |  | Modified HD12-6.8-AG | +12VDC
TB12 - On optical table near dewar
Device | Terminal | Wire | TB12 | TB12 | Wire | Terminal | Device | Notes
TB12 | 3 | Resistor | 1 | 2 | Wht | L | Cable 29 : Connector J4 | 10K
TB12 | 1 | Resistor | 3 | 4 | Grn | K | Cable 29 : Connector J4 | for Thermistor T3 (table-CCD)
Thermistor T3 (table-CCD) |  | Wht | 3 | 4
Thermistor T3 (table-CCD) |  | Wht | 5 | 6 | Yel | W | Cable 29 : Connector J4
Focus rotary transducer | 2 | Org/Blk | 7 | 8 | Org | V | Cable 29 : Connector J4
Focus rotary transducer | 3 | Blk | 9 | 10 | Blk | a | Cable 29 : Connector J4
Focus rotary transducer | 1 | Red | 11 | 12 | Red | J | Cable 29 : Connector J4
TB12 | 15 | Jump | 13 | 14 | Shield |  | Cable 29 : Connector J4
TB12 | 13 | Jump | 15 | 16 | Shield |  | Cable 32 : Connector J3
Ground (table) |  | Grn 18 | 15 | 16 |  | X | Cable 32 : Connector J3
KSL061T1Y motor |  | Blk | 17 | 18 | Blk 18 | W | Cable 32 : Connector J3
KSL061T1Y motor |  | Wht | 19 | 20 | Wht 18 | M | Cable 32 : Connector J3
KSL061T1Y motor |  | Red | 21 | 22 | Red 18 | N | Cable 32 : Connector J3
TB13 - On optical table near middle
Device | Terminal | Wire | TB13 | TB13 | Wire | Terminal | Device | Notes
Cable 28 : Connector J5 | Y | Org | 1 | 2 | Ylw |  | Barometric Pressure Sensor | Analog output ground
Cable 28 : Connector J5 | C | Red | 3 | 4 | Red |  | Barometric Pressure Sensor | Trigger +5V
Cable 28 : Connector J5 | P | Blk | 5 | 6 | Blk |  | Barometric Pressure Sensor | 12V supply ground
Cable 28 : Connector J5 | E | Blu | 7 | 8 | Blu |  | Barometric Pressure Sensor | Analog output ground
Cable 23 : Connector J5 | R | Org | 9 | 10 | Org |  | Barometric Pressure Sensor | 12V supply
Cable 23 : Connector J5 | D | Red | 11 | 12
Cable 23 : Connector J5 |  | Blk | 13 | 14 |  |  |  | Common of 24V variable supply
Cable 28 : Connector J5 |  | Shield | 15 | 16 | Shield |  | Cable to T23
Cable 23 : Connector J5 |  | Shield | 15 | 16 | Grn |  | Ground to table
Cable 28 : Connector J5 | G | Wht | 17 | 18 | Wht | 15 | TB23 | T6 (dewar shell) thermistor
|  |  | 19 | 20 | Blk | 13 | TB23 | T6 (dewar shell) ground
Cable 28 : Connector J5 | T | Brn | 19 | 20
Cable 28 : Connector J5 | F | Yel | 21 | 22 | Red | 11 | TB23 | T6 (dewar shell) resistor
Cable 23 : Connector J5 | Z | Grn | 23 | 24 | Wht |  | Thermistor T2 (table-middle)
|  |  | 25 | 26 | Wht |  | Thermistor T2 (table-middle)
Cable 28 : Connector J5 | S | Grn | 25 | 26 | Resistor | 28 | TB13 | for Thermistor T2 (table-middle)
Cable 23 : Connector J5 | X | Brn | 27 | 28 | Resistor | 26 | TB13 | 10K
Cable 23 : Connector J5 | J | Wht | 29 | 30 | Wht |  | Thermistor T4 (air-inside)
|  |  | 31 | 32 | Silver |  | Thermistor T4 (air-inside)
Cable 23 : Connector J5 | a | Blu | 31 | 32 | Resistor | 34 | TB13 | for Thermistor T4 (air-inside)
Cable 23 : Connector J5 | V | Yel | 33 | 34 | Resistor | 32 | TB13 | 30K & 27K (14.33K measured)
TB14 - On optical table near grating
Device | Terminal | Wire | TB14 | TB14 | Wire | Terminal | Device | Notes
TB9 | 13 | Blk | 1 | 2 | Blk |  | TTL PMT | RG/174U Coaxial cable
TB9 | 15 | Shield | 3 | 4 | Shield |  | TTL PMT | RG/174U Coaxial cable
Cable 24 : Connector J4 | R | Wht | 5 | 6 | Resistor | 8 | TB14 | 10K
Cable 27 : Connector J4 | E | Blk | 7 | 8 | Resistor | 6 | TB14 | for Thermistor T1 (table-grating)
|  |  | 7 | 8 | Wht |  | Thermistor T1 (table-grating)
Cable 24 : Connector J4 | D | Red | 9 | 10 | Wht |  | Thermistor T1 (table-grating)
Cable 27 : Connector J4 |  | Shield | 11 | 12
Cable 24 : Connector J4 |  | Shield | 11 | 12 | Grn 18 |  | Ground to table
Cable 27 : Connector J4 | F | Blk | 13 | 14 | Silver |  | Serial PMT | Used with old Sens-Tech PMT | NC
Cable 27 : Connector J4 | T | Red | 15 | 16 | Red |  | Serial PMT | Used with old Sens-Tech PMT | NC
Cable 27 : Connector J4 | U | Org | 17 | 18 | Blu |  | Serial PMT | Used with old Sens-Tech PMT | NC
Cable 27 : Connector J4 | Z | Blu | 19 | 20 | Blu |  | 250 Ohm heating resistor | Mimics PMT heat when PMT off
Cable 27 : Connector J4 | S | Brn | 21 | 22 | Brn |  | 250 Ohm heating resistor | Mimics PMT heat when PMT off
Cable 27 : Connector J4 | G | Grn | 23 | 24 | Red |  | PMT | Used with either old or new PMT
Cable 27 : Connector J4 | H | Yel | 25 | 26 | Blk |  | PMT | Used with either old or new PMT
TB15 - Attached to Hartmann mask assembly
Device | Terminal | Wire | TB15 | TB15 | Wire | Terminal | Device | Notes
KML063F03E stepper motor |  | Wht/Blk | 1 | 2 | Grn 18 | P | Cable 26 : Connector J3
KML063F03E stepper motor |  | Blk | 3 | 4 | Blk 18 | B | Cable 26 : Connector J3
KML063F03E stepper motor |  | Wht/Red | 5 | 6 | Wht 18 | R | Cable 26 : Connector J3
KML063F03E stepper motor |  | Red | 7 | 8 | Red 18 | A | Cable 26 : Connector J3
Hartmann rotary transducer | 3 | Blk | 9 | 10 | Blk | C | Cable 25 : Connector J4
Hartmann rotary transducer | 2 | Org | 11 | 12 | Wht | Y | Cable 25 : Connector J4
Hartmann rotary transducer | 1 | Red | 13 | 14 | Red | P | Cable 25 : Connector J4
TB16 - Inside Slit assembly (split into two groups)
Device | Terminal | Wire | TB16 | TB16 | Wire | Terminal | Device | Notes
J7 | 2 | Brn | 1 | 2 | Red |  | Shutter
J7 | 1 | Gry | 3 | 4 | Blk |  | Shutter
J7 | 9 | Brn | 5 | 6 | Blk |  | Diffuser Shutter
J7 | 10 | Gry | 7 | 8 | Red |  | Diffuser Shutter
Slit motor |  | Brn | 9 | 10 | Brn | 11 | J7
Slit motor |  | Gry | 11 | 12 | Gry | 12 | J7
Slit linear transducer |  | Red | 13 | 14 | Red | 13 | J7
Slit linear transducer |  | Blk | 15 | 16 | Blk | 14 | J7
Slit linear transducer |  | Org | 17 | 18 | Org | 15 | J7
TB17 - Inside pre-slit box, away from connectors
Device | Terminal | Wire | TB17 | TB17 | Wire | Terminal | Device | Notes
New Focus flipper |  | Red | 1 | 2 | Red | 15 | J1
New Focus flipper |  | Wht | 3 | 4 | Wht | 14 | J1
New Focus flipper |  | Shield | 5 | 6 | Blk | 1 | J1
Guider image shutter |  | Blu/Red | 7 | 8 | Brn | 21 | J1
Guider image shutter |  | Red/Blu | 9 | 10 | Blk | 24 | J1
Iodine cell motor |  | Brn | 11 | 12 | Brn | 17 | J1
Iodine cell motor |  | Gry | 13 | 14 | Gry | 16 | J1
Limit switch 2 |  | Blk/Blu | 15 | 16 | Blk/Blu | 23 | J1
Limit switch 2 |  | Blu/Blk | 17 | 18 | Blu/Blk | 22 | J1
|  |  | 19 | 20
|  |  | 21 | 22
|  |  | 21 | 22
TB18 - Inside pre-slit box, near connectors
Device | Terminal | Wire | TB18 | TB18 | Wire | Terminal | Device | Notes
Iodine cell heat wrap |  | Wht | 1 | 2 | Wht 18 | C | J2
Iodine cell heat wrap |  | Wht | 3 | 4 | Blk 18 | A | J2
Iodine cell thermocouple |  | Yel | 5 | 6 | Yel | 5 | J1
Iodine cell thermocouple |  | Red | 7 | 8 | Red | 6 | J1
|  |  | 9 | 10 | Blk/Org | 7 | J1
Iodine cell temp. sensor 2 |  | Org/Blk | 11 | 12 | Org/Blk | 8 | J1
Limit switch 1 |  | Blk/Blu | 13 | 14 | Blk/Blu | 9 | J1
Limit switch 1 |  | Blu/Blk | 15 | 16 | Blu/Blk | 10 | J1
Iodine linear transducer |  | Org | 17 | 18 | Org | 11 | J1
Iodine linear transducer |  | Red | 19 | 20 | Red | 12 | J1
Iodine linear transducer |  | Blk | 21 | 22 | Blk | 13 | J1
Iodine cell temp. sensor 2 |  | Blk/Org | 21 | 22
TB19 - Outside pre-slit box, on right side
Device | Terminal | Wire | TB19 | TB19 | Wire | Terminal | Device | Notes
QTH lamp |  | Blk | 1 | 2 | Blk |  | Cable 6
QTH lamp |  | Red | 3 | 4 | Red |  | Cable 6
ThAr lamp |  | Blu | 5 | 6 | Blu |  | Cable 6
ThAr lamp |  | Red | 7 | 8 | Vio |  | Cable 6
TB20 - Inside wooden, insulated heating system box
Device | Terminal | Wire | TB20 | TB20 | Wire | Terminal | Device | Notes
Thermistor T7 (glycol supply) |  | Wht | 1 | 2 | Blu |  | Cable 18
Thermistor T7 (glycol supply) |  | Shield | 3 | 4 | Org |  | Cable 18
TB20 | 5 | Resistor | 3 | 4
TB20 | 3 | Resistor | 5 | 6 | Yel |  | Cable 18 | for Thermistor T7 (glycol supply)
Pump |  | Blk 14 | 7 | 8 | Blk 14 |  | Cable 17
Pump |  | Red 14 | 9 | 10 | Red 14 |  | Cable 17
Flowmeter |  | Wht | 11 | 12 | Wht |  | Cable 18
Flowmeter |  | Grn | 13 | 14 | Grn |  | Cable 18
Flowmeter |  | Brn | 15 | 16 | Brn |  | Cable 18
Thermistor T8 (glycol return) |  | Wht | 17 | 18 | Wht |  | Cable 19
Thermistor T8 (glycol return) |  | Shield | 19 | 20 | Blk |  | Cable 19
TB20 | 21 | Resistor | 19 | 20
TB20 | 19 | Resistor | 21 | 22 | Grn |  | Cable 19 | for Thermistor T8 (glycol return)
|  |  | 23 | 24 | Grn 14 |  | Cable 17
|  |  | 25 | 26 | Brn 14 |  | Cable 17
|  |  | 27 | 28 | Red
|  |  | 29 | 30 | Blk
|  |  | 31 | 32
|  |  | 33 | 34
|  |  | 33 | 34
AC Tape 75W |  | Blk | 35 | 36 | Blk |  |  | AC Live
AC Tape 75W |  | Blk | 37 | 38 | Grn |  |  | AC Live
AC Tape 75W |  | Blk | 39 | 40 | Red |  |  | AC Live
AC Tape 75W |  | Blk | 41 | 42 | Brn |  |  | AC Live
AC Tape 75W |  | Blk | 43 | 44 | Wht |  |  | AC Neutral
AC Tape 75W |  | Blk | 43 | 44 | Jump | 46 | TB20 | AC Neutral
AC Tape 75W |  | Blk | 45 | 46 | Jump | 44 | TB20 | AC Neutral
AC Tape 75W |  | Blk | 45 | 46 | Blu |  |  | AC Neutral
TB21 – Attached to cart underneath End 1
Device | Terminal | Wire | TB21 | TB21 | Wire | Terminal | Device | Notes
TB8 | 51 | Blk | 1 | 2 | Wht |  | Thermistor T5 (air-outside)
TB8 | 25 | Red | 3 | 4 | Shield |  | Thermistor T5 (air-outside)
|  |  | 3 | 4 | Resistor
TB8 | 49 | Wht | 5 | 6 | Resistor |  |  | for Thermistor T5 (air-outside)
TB22 – Attached to floor outer face underneath End 2
Device | Terminal | Wire | TB22 | TB22 | Wire | Terminal | Device | Notes
Electronics box mount plate |  | Grn 18 | 1 | 2 | Grn 18 |  | Insulation side 2 pipe clamp | Ground
TB22 | 3 | Jump | 1 | 2 | Grn 18 |  | Insulation floor pipe clamp | Ground
TB22 | 1 | Jump | 3 | 4 |  |  |  | Ground
|  |  | 3 | 4 |  |  |  | Ground
TB23 – Attached to dewar mount
Device | Terminal | Wire | TB23 | TB23 | Wire | Terminal | Device | Notes
|  |  | 1 | 2 | Red & Blk |  | Heat tape on dewar side
Cable to Hoffman box |  | Red 18 | 1 | 2 | Blk & Blk |  | Heat tape on dewar side | 16VAC nominal, adjusted with pot
Cable to Hoffman box |  | Blk 18 | 3 | 4 | Red & Blk |  | Heat tape on dewar side
|  |  | 3 | 4 | Blk & Blk |  | Heat tape on dewar side
|  |  | 5 | 6
Cable to Hoffman box |  | Wht 18 | 5 | 6 |  |  | Heat tape on lN2 fill tube extension | 16VAC nominal, adjusted with pot
Cable to Hoffman box |  | Blk 18 | 7 | 8 |  |  | Heat tape on lN2 fill tube extension
|  |  | 7 | 8
Cable to Hoffman box |  | Red 18 | 9 | 10 | Brn |  | Heat tape on dewar end
Cable to Hoffman box |  | Blk 18 | 11 | 12 | Blu |  | Heat tape on dewar end
|  |  | 13 | 14
TB13 |  | Red | 15 | 16 | Resistors |  | T6 (dewar shell) resistor | 30K & 27 K (14.07K measured)
TB13 |  | Blk | 17 | 18 | Resistors |  | T6 (dewar shell) ground
|  |  | 17 | 18 | Blk |  | T6 (dewar shell) thermistor
TB13 |  | Wht | 19 | 20 | Wht |  | T6 (dewar shell) thermistor
TB24 – Inside Hoffman box below stepper motor controller
Device | Terminal | Wire | TB24 | TB24 | Wire | Terminal | Device | Notes
Hartmann stepper motor, via J3-B | A+ | Blk 18 | 1 | 2 | Blk | A+ | STP-DRV-4035
Hartmann stepper motor, via J3-P | A- | Grn 18 | 3 | 4 | Grn | A- | STP-DRV-4035
Hartmann stepper motor, via J3-A | B+ | Red 18 | 5 | 6 | Red | B+ | STP-DRV-4035
Hartmann stepper motor, via J3-R | B- | Wht 18 | 7 | 8 | Wht | B- | STP-DRV-4035
TB25 – On air heating tube on CCD end
Device | Terminal | Wire | TB24 | TB24 | Wire | Terminal | Device | Notes
|  |  | 1 | 2
Air Heat Fan |  | Blk | 1 | 2 | Blk | 18 | TB6 | Ground
Air Heat Fan |  | Red | 3 | 4 | Red | C1 | F2-08TRS Relay - base 0, slot 5 | 12V through 40ohm resistor
|  |  | 5 | 6 | Blk |  |  | Not connected
|  |  | 7 | 8 | Grn |  |  | Not connected
|  |  | 9 | 10
Heat Resistor Pair 1 |  | Blk | 11 | 12 | Blk | 2 | F2-08TA AC out
Heat Resistor Pair 1 |  | Blu | 13 | 14 | Blu | 3 | TB11
Heat Resistor Pair 2 |  | Blk | 15 | 16 | Blk | 3 | F2-08TA AC out
Heat Resistor Pair 2 |  | Wht | 17 | 18 | Wht | 3 | TB11
TB26 – On air heating tube on grating end
Device | Terminal | Wire | TB24 | TB24 | Wire | Terminal | Device | Notes
TB25 |  | Blk | 1 | 2 | Blk/Grn | 15 | Heat Resistor Pair 2 | AC Hot
TB25 |  | Wht | 3 | 4 | Wht | 17 | Heat Resistor Pair 2
|  |  | 5 | 6
|  |  | 7 | 8

## PLC
PLC Terminal Wiring
Base | Slot | Module | Terminal | Color | Connected to | Terminal | Notes | Input/Output
0 | 0 | D2-CTRINT Counter |  |  |  |  | Consumes 8 Input and 8 Output points | X0-X7,Y0-Y7
0 | 0 | D2-CTRINT Counter | C2 | Red | TB6 | 24 | to STP-DRV-4035 Dir+/Step+
0 | 0 | D2-CTRINT Counter | 3 | Wht/Org | TB6 | 12 | to STP-DRV-4035 Step-
0 | 0 | D2-CTRINT Counter | 4 | Org/Wht | TB6 | 14 | to STP-DRV-4035 Dir-
0 | 0 | D2-CTRINT Counter | C4 | Blk | TB6 | 16 | Common (Ground)
0 | 1 | H2-ECOM100 Communication |  |  |  |  | Consumes no Input/Output points | none
0 | 1 | H2-ECOM100 Communication | 10Base-T | Blu 10T | Back panel ethernet connector | 10Base-T
0 | 2 | F2-CP128 Coprocessor |  |  |  |  | Consumes no Input/Output points | none
0 | 2 | F2-CP128 Coprocessor | Port1/3 | Blk Serial | PC (port 1), PMT (Port 3) | Splitter
0 | 2 | F2-CP128 Coprocessor | Port 2 | Blk Serial | SMC100 | RS232C
0 | 3 | F2-08TRS Relay |  |  |  |  | Consumes 8 output points | Y10-Y17
0 | 3 | F2-08TRS Relay | NO0 | Org/Wht | F2-08TRS Relay - base 0, slot 3 | NO7
0 | 3 | F2-08TRS Relay | C0 | Org/Blk | TB6 | 32 | to Pur/Or slit speed pot | Y10
0 | 3 | F2-08TRS Relay | NC0 | Org/Wht | F2-08TRS Relay - base 0, slot 3 | NO6
0 | 3 | F2-08TRS Relay | C3 | Wht 18 | TB6 | 34 | to +5V Sola PS | Y13
0 | 3 | F2-08TRS Relay | NO3 | Red 18 | TB6 | 40 | to Quartz lamp
0 | 3 | F2-08TRS Relay | NC6 | Blk | TB6 & F2-08TRS Relay - base 0, slot 3 | 18, NC7 | Common
0 | 3 | F2-08TRS Relay | C6 | Brn/Vio | TB6 | 46 | to Slit Motor Brown | Y16
0 | 3 | F2-08TRS Relay | NO6 | Org/Wht | F2-08TRS Relay - base 0, slot 3 | NC0
0 | 3 | F2-08TRS Relay | NC7 | Blk | F2-08TRS Relay - base 0, slot 3 | NC6
0 | 3 | F2-08TRS Relay | C7 | Vio/Brn | TB6 | 48 | to Slit Motor Gray | Y17
0 | 3 | F2-08TRS Relay | NO7 | Org/Wht | F2-08TRS Relay - base 0, slot 3 | NO0
0 | 4 | F2-08TRS Relay |  |  |  |  | Consumes 8 output points | Y20-Y27
0 | 4 | F2-08TRS Relay | NO0 | Org | F2-08TRS Relay - base 0, slot 4 | NO7
0 | 4 | F2-08TRS Relay | C0 | Org/Rd | TB6 | 36 | to Iodine carriage speed pot | Y20
0 | 4 | F2-08TRS Relay | NC0 | Org | F2-08TRS Relay - base 0, slot 4 | NO6
0 | 4 | F2-08TRS Relay | C1 | Org/Blk | TB6 | 26 | +5V
0 | 4 | F2-08TRS Relay | NO1 | Grn/Yel | TB6 | 8 | to Flipper controller TTL IN | Y21
0 | 4 | F2-08TRS Relay | C2 | Red | TB6 | 26 | +5V
0 | 4 | F2-08TRS Relay | NO2 | Yel/Brn | TB5 | 38 | to HV supply & Ground w/ pull-down 1K5 | Y22
0 | 4 | F2-08TRS Relay | C3 | Red | TB6 | 38 | 22ohm resist to TB6 terminal 41 | Y23
0 | 4 | F2-08TRS Relay | NO3 | Red | TB6 | 36 | to Iodine carriage speed pot
0 | 4 | F2-08TRS Relay | C5 | Yel 18 | TB5 | 22 | +24V
0 | 4 | F2-08TRS Relay | NO5 | Yel 18 | TB6 | 30 | to STP-DRV-4035 VDC+ | Y25
0 | 4 | F2-08TRS Relay | NC6 | Blk/Red | TB6 & F2-08TRS Relay - base 0, slot 4 | 18, NC7 | Common
0 | 4 | F2-08TRS Relay | C6 | Blu/Vio | TB6 | 50 | to Iodine carriage motor brown | Y26
0 | 4 | F2-08TRS Relay | NO6 | Org | F2-08TRS Relay - base 0, slot 4 | NC0
0 | 4 | F2-08TRS Relay | NC7 | Blk/Red | F2-08TRS Relay - base 0, slot 4 | NC6
0 | 4 | F2-08TRS Relay | C7 | Vio/Blu | TB6 | 52 | to Iodine carriage motor gray | Y27
0 | 4 | F2-08TRS Relay | NO7 | Org | F2-08TRS Relay - base 0, slot 4 | NO0
0 | 5 | F2-08TRS Relay |  |  |  |  | Consumes 8 output points | Y30-Y37
0 | 5 | F2-08TRS Relay | NO0 | Org | F2-08TRS Relay - base 0, slot 5 | NO7
0 | 5 | F2-08TRS Relay | C0 | Org/Blu | TB6 | 62 | through series resistor to +12V | Y30
0 | 5 | F2-08TRS Relay | NC0 | Org | F2-08TRS Relay - base 0, slot 5 | NO6
0 | 5 | F2-08TRS Relay | C3 | Blue,Blu | TB9 & F2-08TRS Relay | 35, C5 | Output from transformer (20VAC measured)
0 | 5 | F2-08TRS Relay | NO3 | Blu | TB7 | 13 | PMT heater AC hot | Y33
0 | 5 | F2-08TRS Relay | C4 | Blk |  |  | Ground for guider shutter
0 | 5 | F2-08TRS Relay | NO4 | Blu |  |  | to guider shutter | Y34
0 | 5 | F2-08TRS Relay | C5 | Blu | F2-08TRS Relay | C3 | Output from transformer (20VAC measured)
0 | 5 | F2-08TRS Relay | NO5 | Blu | TB9 | 30 | Side dewar heater AC hot | Y35
0 | 5 | F2-08TRS Relay | NC6 | Blk/Gry | TB6 & F2-08TRS Relay - base 0, slot 5 | 20, NC7 | Common
0 | 5 | F2-08TRS Relay | C6 | Brn/Vio | TB6 | 54 | to Shutter gray | Y36
0 | 5 | F2-08TRS Relay | NO6 | Org | F2-08TRS Relay - base 0, slot 5 | NC0
0 | 5 | F2-08TRS Relay | NC7 | Blk/Gry | F2-08TRS Relay - base 0, slot 5 | NC6
0 | 5 | F2-08TRS Relay | C7 | Vio/Brn | TB6 | 56 | to Shutter brown | Y37
0 | 5 | F2-08TRS Relay | NO7 | Org | F2-08TRS Relay - base 0, slot 5 | NO0
0 | 6 | F2-08TRS Relay |  |  |  |  | Consumes 8 output points | Y40-Y47
0 | 6 | F2-08TRS Relay | NO0 | Org | F2-08TRS Relay - base 0, slot 6 | NO7
0 | 6 | F2-08TRS Relay | C0 | Org/Blu | TB6 | 66 | through series resistor to +12V | Y40
0 | 6 | F2-08TRS Relay | NC0 | Org | F2-08TRS Relay - base 0, slot 6 | NO6
0 | 6 | F2-08TRS Relay | NO1 | Red | TB4 | 1 | 40 Ohm resistor in line to 12VDC for air heat fan | Y41
0 | 6 | F2-08TRS Relay | C1 | Red
0 | 6 | F2-08TRS Relay | C3 | Wht | TB5 | 16 | Interlock for Ion pump | Y43
0 | 6 | F2-08TRS Relay | NO3 | Grn | TB5 | 18
0 | 6 | F2-08TRS Relay | C4 | Red | TB4 | 3 | 5V power to PMT package | Y44
0 | 6 | F2-08TRS Relay | NO4 | Red/Wht | TB7 | 1
0 | 6 | F2-08TRS Relay | NC6 | Blk/Gry | TB6 & F2-08TRS Relay - base 0, slot 6 | 28, NC7 | Common
0 | 6 | F2-08TRS Relay | C6 | Blk/Vio | TB6 | 58 | to Diffuser shutter gray | Y46
0 | 6 | F2-08TRS Relay | NO6 | Org | F2-08TRS Relay - base 0, slot 6 | NC0
0 | 6 | F2-08TRS Relay | NC7 | Blk/Gry | F2-08TRS Relay - base 0, slot 6 | NC6
0 | 6 | F2-08TRS Relay | C7 | Vio/Blu | TB6 | 60 | to Diffuser shutter brown | Y47
0 | 6 | F2-08TRS Relay | NO7 | Org | F2-08TRS Relay - base 0, slot 6 | NO0
0 | 7 | H2-CTRIO High speed Counter
0 | 7 | H2-CTRIO High speed Counter | 1A | Brn | Board 2 | D | output from PMT, thru 4x 74ALS74 & transistor
0 | 7 | H2-CTRIO High speed Counter | 1M | Org | TB4 | 1 | +12V
1 | 0 | F2-08TA AC out |  |  |  |  | Consumes 8 output points | Y50-Y57
1 | 0 | F2-08TA AC out | C0-3 | Brn 18 | DL205 Base power terminal | AC | AC line
1 | 0 | F2-08TA AC out | 0 | Red 18 | TB5 | 10 | to Focus motor black & Board | Y50
1 | 0 | F2-08TA AC out | 1 | Blk 18 | TB5 | 12 | to Focus motor red & Board | Y51
1 | 0 | F2-08TA AC out | 2 | Blk | TB25 | 12 | Air heat resistor pair 1 | Y52
1 | 0 | F2-08TA AC out | 3 | Blk | TB25 | 14 | Air heat resistor pair 2 | Y53
1 | 0 | F2-08TA AC out | C4-7 | Brn 18 | TB11 | 1
1 | 0 | F2-08TA AC out | 4 | Brn 18 | TB11 | 7 | 75W AC Heat tape 1 | Y54
1 | 0 | F2-08TA AC out | 5 | Red 18 | TB11 | 9 | 75W AC Heat tape 2 | Y55
1 | 0 | F2-08TA AC out | 6 | Grn 18 | TB11 | 11 | 75W AC Heat tape 3 | Y56
1 | 0 | F2-08TA AC out | 7 | Blk 18 | TB11 | 13 | 75W AC Heat tape 4 | Y57
1 | 1 | F2-08DA2 Analog output |  |  |  |  | Consumes 16 output points | Y60-Y77
1 | 1 | F2-08DA2 Analog output | 0V | Blk | TB5 | 30 | Common
1 | 1 | F2-08DA2 Analog output | +24V | Yel/Org | TB5 | 22 | +24V
1 | 1 | F2-08DA2 Analog output | +V1 | Wht | TB5 | 56 | +5V trigger for CCD charge shuffle
1 | 1 | F2-08DA2 Analog output | +V2 | Red/Grn | Board | 29 | +5V enable for T1-T6
1 | 1 | F2-08DA2 Analog output | +V3 | Red/Brn | Board | 30 | +5V enable for T7-T8
1 | 1 | F2-08DA2 Analog output | +V4 | Org/Yel | Terminal for control voltage |  | Modified HD12-6.8-AG
1 | 1 | F2-08DA2 Analog output | +V5 | Yel/Blk | Terminal for control voltage |  | Modified HD24-4.8-AG
1 | 2 | D2-32ND3-2 |  |  |  |  | Consumes 32 input points from ZL-CM40 | X10-X47
N/A | N/A | ZIPlink ZL-CM40 | A0 | Red |  |  | CCD saddlebag | X10
N/A | N/A | ZIPlink ZL-CM40 | CI lo | Blk | TB6 | 19 | Common
N/A | N/A | ZIPlink ZL-CM40 | B0 | Wht | SMC100 | GPIO 13 | “In Motion” | X20
N/A | N/A | ZIPlink ZL-CM40 | B1 | Grn | SMC100 | GPIO 14 | “Not Referenced” | X21
N/A | N/A | ZIPlink ZL-CM40 | CII lo | Red | TB7 | 17 | +5V (Patricio's "trick")
N/A | N/A | ZIPlink ZL-CM40 | C0 | Red | Limit switch |  | Reverse limit | X30
N/A | N/A | ZIPlink ZL-CM40 | C1 | Grn | Limit switch |  | Forward limit | X31
N/A | N/A | ZIPlink ZL-CM40 | CIII lo | Red | TB5 | 1 | +5V
1 | 3 | F2-08AD-2 Analog input |  |  |  |  | Consumes 16 input points | X50-X67
1 | 3 | F2-08AD-2 Analog input | 0V | Blk | TB5 | 32 | Common
1 | 3 | F2-08AD-2 Analog input | +24V | Yel/Brn | TB5 | 24 | +24V
1 | 3 | F2-08AD-2 Analog input | CH1+ | Brn/Yel | TB5 | 40 | to slit linear transducer orange | X50
1 | 3 | F2-08AD-2 Analog input | CH2+ | Yel/Brn | TB5 | 44 | to focus rotary transducer orange | X51
1 | 3 | F2-08AD-2 Analog input | CH3+ | Gry/Blk | TB5 | 46 | to HC2024R green & ThAr brightness pot | X52
1 | 3 | F2-08AD-2 Analog input | CH4+ | Wht/Brn | TB9 | 5 | Ion pump | X53
1 | 3 | F2-08AD-2 Analog input | CH5+ | Yel/Org | Board | pin 48 | for grating vacuum gauge | X54
1 | 3 | F2-08AD-2 Analog input | CH6+ | Gry/Yel | Board | pin 50 | for dewar vacuum gauge | X55
1 | 3 | F2-08AD-2 Analog input | CH7+ | Yel/Brn | TB5 | 48 | to Iodine linear transducer orange | X56
1 | 3 | F2-08AD-2 Analog input | CH8+ | Pur/Gry | TB5 | 52 | to Hartmann rotary transducer orange | X57
1 | 4 | F2-08AD-2 Analog input |  |  |  |  | Consumes 16 input points | X70-X107
1 | 4 | F2-08AD-2 Analog input | 0V | Blk | TB5 | 36 | Common
1 | 4 | F2-08AD-2 Analog input | +24V | Yel/Brn | TB5 | 24 | +24V
1 | 4 | F2-08AD-2 Analog input | CH1+ | Blu/Vio | Board | pin 28 | temperature T1 (table near grating) | X70
1 | 4 | F2-08AD-2 Analog input | CH2+ | Vio/Blu | Board | pin 25 | temperature T2 (table in middle) | X71
1 | 4 | F2-08AD-2 Analog input | CH3+ | Gry/Vio | Board | pin 22 | temperature T3 (table near CCD) | X72
1 | 4 | F2-08AD-2 Analog input | CH4+ | Vio/Gry | Board | pin 19 | temperature T4 (spectrograph air) | X73
1 | 4 | F2-08AD-2 Analog input | CH5+ | Blu/Vio | Board | pin 16 | temperature T5 (dome air) | X74
1 | 4 | F2-08AD-2 Analog input | CH6+ | Vio/Blu | Board | pin 13 | temperature T6 (TBD) | X75
1 | 4 | F2-08AD-2 Analog input | CH7+ | Gry/Vio | Board | pin 10 | temperature T7 (liquid out) | X76
1 | 4 | F2-08AD-2 Analog input | CH8+ | Vio/Gry | Board | pin 7 | temperature T8 (liquid in) | X77
1 | 5 | F2-08AD-2 Analog input |  |  |  |  | Consumes 16 input points | X110-X127
1 | 5 | F2-08AD-2 Analog input | 0V | Blk |  | 32 | Common
1 | 5 | F2-08AD-2 Analog input | +24V | Yel/Brn |  | 21 | +24V
1 | 5 | F2-08AD-2 Analog input | CH1+ | Org | TB17 | 1 | to Iodine cell temperature sensor | X110
1 | 5 | F2-08AD-2 Analog input | CH2+ | Org | Connector J5 | Y | Barometric pressure | X111
1 | 5 | F2-08AD-2 Analog input | CH3+ | Vio/Brn | Board | pin 46 | to flowmeter totalizer | X112
1 | 5 | F2-08AD-2 Analog input | CH4+ | Brn/Wht | Board | Pin 41 | +2.5V from V divider + follower | X113
1 | 5 | F2-08AD-2 Analog input | CH5+
1 | 5 | F2-08AD-2 Analog input | CH6+
1 | 5 | F2-08AD-2 Analog input | CH7+
1 | 5 | F2-08AD-2 Analog input | CH8+